## Slide 1
HD03 Hemodialysis Monitor Training
Delivered Flow Measurement

## Slide 2: Delivered Flow
Prerequisite includes reviewing the HD03 Hemodialysis Monitor PowerPoint Training Introduction To be used with Delivered Flow, Recirculation, and Access Flow
Detailed Delivered Flow Measurement Procedure Steps
Clinical interpretation for Catheters and AV Fistula of Grafts

## Slide 3: Set up for Delivered Flow Measurement
Utilizes the steps detailed in the HD03 Hemodialysis Monitor PowerPoint Training Introduction To be used with Delivered Flow, Recirculation, and Access Flow to select the proper bloodlines (tubing), sensor placement, basic troubleshooting, and disinfection procedures.

## Slide 4: Proper Sensor Placement on Bloodlines (Tubing)
Sensor Placement
Properly placed Arterial & Venous Sensors on the bloodlines (Tubing)

## Slide 5: Choose Delivered Flow Protocol

## Slide 6: Delivered Flow Measurement
Does Your Dialysis Machine Use Compensated Blood Flow Rate?
(Examples Fresenius 5008 or Baxter Phoenix)
If Yes:
If No:
If the compensated blood flow rate is used to set the patient prescription, then enter the compensated blood flow rate into the Transonic pump setting box
If the dialysis machine blood pump setting is used to set the patient prescription, then enter the dialysis machine blood pump setting into the Transonic pump setting box
Enter the dialysis machine blood pump setting into the Transonic pump setting box

## Slide 7: What is Compensated Blood Flow Rate?
Pump Speed is the speed set on the dialysis machine blood pump in mL/min
Compensated Blood Flow Rate is the calculated Blood Flow Rate adjusted to the current Negative Arterial Pressure
The Physician prescribes the Blood Flow Rate (BFR) as part of the hemodialysis prescription. Check your dialysis facility policy & procedure for guidance on the Blood Flow Rate that is used to meet the prescription. Some facilities increase the Pump Speed as needed to achieve the Compensated Blood Flow Rate that matches the prescribed BFR.

## Slide 8: Delivered Flow Measurement
Procedure Steps
Enter the Pump Setting Speed in mL/min using the number keypad
Then touch the Measure icon
The Delivered Flow will display
If the Delivered Flow is different than the Pump Setting, the % difference will post
# 1
# 3
# 4
# 2

## Slide 9: Delivered Flow Measurement
Result Interpretation
Within Range
The Delivered Flow within 10% of the Pump Setting is considered within range. The results box is green.
The Delivered Flow is more accurate as it is the actual blood flow rate inside the circuit.

## Slide 10: Delivered Flow Measurement
Result Interpretation
Out of Range
If the Delivered Flow is not within 10% of the Pump Setting, the percentage difference will display in the result box in red.
If the Delivered Flow is lower than the Pump Setting, this means the patient is getting less dialysis than prescribed.
Check your facility’s policy on whether to adjust the Pump Setting to improve the patient’s Delivered Flow.

## Slide 11: Delivered Flow Measurement
Result Interpretation
Out of Range
Check for the proper selection of the bloodlines (tubing) used for the measurement – wrong tubing selection will impact the measurement accuracy.
Check for kinking of bloodlines as this is also a hemolysis risk.
Needle type and gauge may impact the Delivered Flow.
AVG/AVF condition and Access Flow may impact Delivered Flow. Recirculation and Access Flow results should be obtained for full assessment.
Check the machine blood pump calibration.

## Slide 12: Delivered Flow Measurement on HD Machine with Compensated Blood Flow
Results Interpretation
The Delivered Flow can be higher than the Pump Setting with dialysis machines that utilized a compensated blood flow.

## Slide 13: Delivered Flow Results Interpretation By Access Type
Catheters
AV Graft or AV Fistula
If the Delivered Flow is greater than 10% or more to the Pump Setting, then the patient is not achieving the targeted dose of dialysis
Check for kinks in the catheter or bloodlines as this is also a hemolysis risk
Utilize the Optimization of Hemodialysis Adequacy for Catheter Connection Configuration with the Transonic Flow-QC® Hemodialysis Monitor algorithm to measure Delivered Flow and Recirculation to guide an on-the-spot correction.
Check for kinking of bloodlines as this is also a hemolysis risk
Check for miss match of needle gauge for the prescribed Blood Flow Rate (example 16-gauge needles and 450 BFR)
Excessively negative pre-pump Arterial Pressure can limit the Delivered Flow and is most likely due to a very low Access Flow

## Slide 14: Optimization of Hemodialysis Catheter Adequacy Algorithm
Dialysis Machine Without Compensated Blood Flow Rate
Dialysis Machine With Compensated Blood Flow Rate
Optimization of Hemodialysis Catheter Adequacy Algorithm without Compensated Blood Flow Rate
Optimization of Hemodialysis Catheter Adequacy Algorithm with Compensated Blood Flow Rate

## Slide 15: Transonic Tools at www.Transonic.com

## Slide 16: Transonic HD03 Tools www.Transonic.com
Customer Updates & Training Web Portal
Software Updates
Training Materials
Training Videos
Training Checklists
Operator’s Manuals

## Slide 17: Customer Login to Access the Partner Portal
You can click on any of the three buttons to access the Partner Portal
Top right of the home page
Click on Contact Us then select Partner Portal
Click on Software Updates on the bottom of the home page

## Slide 18: Complete the Form to Create Your Account
Select HD03 from the dropdown box
Use a facility email address if your facility has a general email address.
Notifications about your HD03 will be sent to the email used to register.

## Slide 19: Email Confirmation

## Slide 20: Partner Portal Home Page
Click on HD03 Customer

## Slide 21: Partner Portal Home Page
You can save the password on your computer, so it will auto populate and go to the Partner Portal.

## Slide 22: Partner Portal
Scroll down on the page to find the Tubing List, Software Updates & Manuals

## Slide 23: Additional Questions?
Please refer to the HD03 Operator’s Manual, Quick Reference Guides, or contact:
Transonic Tech Support 800-353-3569 (US and Canada)
www.Transonic.com